## AGW (Bombas)
Code | Level | Class | Text | Short Text | Level2 | Level3
120 | MediumPriorityAlarm | Technical | Sequence number incorrect (lower) | Sequence number incorrect (lower) | 0
121 | MediumPriorityAlarm | Technical | Anomaly on AGW. Infusion parameters are not changed. Restart AGW. | Anomaly on AGW. Infusion parameters are not changed. Restart AGW. | 0
122 | MediumPriorityAlarm | Technical | AGW Version not supported: {0} {1} | AGW Version not supported: {0} {1} | 0
123 | MediumPriorityAlarm | Technical | Battery level low. {0} left | Battery level low. {0} left | 0
200 | ClinicalEvent | Physiological | Bolus | Bolus | 0
255 | UserEvent | Technical | Titration Event | Titration Event | 0
256 | UserEvent | Technical | Pump disconnected {0} {1} | Pump disconnected {0} {1} | 1
257 | UserEvent | Technical | Pump connected {0} {1} | Pump connected {0} {1} | 1
260 | UserEvent | Technical | Barcode read: {0} | Barcode read: {0} | 0
ENT 01 | HighPriorityAlarm | Technical | AC POWER FAIL | AC POWER FAIL | 0
ENT 02 | HighPriorityAlarm | Technical | DRIVE DISENGAGED | DRIVE DISENGAGED | 0
ENT 03 | HighPriorityAlarm | Technical | BATTERY EMPTY | BATTERY EMPTY | 1
ENT 04 | HighPriorityAlarm | Physiological | END OF DELIVERY | END OF DELIVERY | 1
ENT 05 | HighPriorityAlarm | Physiological | END OF DELIVERY | END OF DELIVERY | 1
ENT 06 | HighPriorityAlarm | Physiological | END OF DELIVERY | END OF DELIVERY | 1
ENT 07 | HighPriorityAlarm | Technical | BATTERY LOW | BATTERY LOW | 1
ENT 08 | HighPriorityAlarm | Physiological | NEAR END OF DELIVERY | NEAR END OF DELIVERY | 0 | *
ENT 09 | HighPriorityAlarm | Physiological | OCCLUSION | OCCLUSION | 1
ENT 10 | HighPriorityAlarm | Physiological | CHECK SYRINGE | CHECK SYRINGE | 1
ENT 11 | HighPriorityAlarm | Physiological | TITRATION NOT CONFIRMED | TITRATION NOT CONFIRMED | 0
ENT 12 | HighPriorityAlarm | Physiological | VTBI DONE | VTBI DONE | 0
ENT 13 | HighPriorityAlarm | Physiological | VTBI DONE | VTBI DONE | 0
ENT 14 | HighPriorityAlarm | Physiological | VTBI DONE - STOP | VTBI DONE - STOP | 0
ENT 15 | HighPriorityAlarm | Technical | Attention (with " 3 beeps") | Attention (with " 3 beeps") | 0
ENT 17 | HighPriorityAlarm | Technical | COMMUNICATIONS TIMEOUT | COMMUNICATIONS TIMEOUT | 0
ENT 18 | HighPriorityAlarm | Physiological | INFUSION LINE DISCONNECTION | INFUSION LINE DISCONNECTION | 1
ENT 19 | HighPriorityAlarm | Physiological | EXCESSIVE RATE | EXCESSIVE RATE | 0
GW 01 | HighPriorityAlarm | Technical | bAt (INTERNAL BATTERY DEPLETED) | bAt (INTERNAL BATTERY DEPLETED) | 0
GW 02 | HighPriorityAlarm | Physiological | Air OCCL (AIR-IN-LINE - UPSTREAM OCCLUSION) | Air OCCL (AIR-IN-LINE - UPSTREAM OCCLUSION) | 0
GW 03 | HighPriorityAlarm | Physiological | HI PrES (DOWNSTREAM OCCLUSION. A blockage has occurred downstream) | HI PrES (DOWNSTREAM OCCLUSION. A blockage has occurred downstream) | 0
GW 04 | HighPriorityAlarm | Physiological | door (DOOR OPEN. Door was opened during an infusion) | door (DOOR OPEN. Door was opened during an infusion) | 0
GW 05 | HighPriorityAlarm | Technical | bAd SEt (INCORRECT INFUSION SET. Set incorrectly loaded or set worn) | bAd SEt (INCORRECT INFUSION SET. Set incorrectly loaded or set worn) | 0
GW 06 | HighPriorityAlarm | Physiological | End (Finished pre-set volume to be infused.) | End (Finished pre-set volume to be infused.) | 0
GW 07 | HighPriorityAlarm | Physiological | FLo SEnS (FLOW SENSOR CONNECTION ERROR) | FLo SEnS (FLOW SENSOR CONNECTION ERROR) | 0
GW 08 | HighPriorityAlarm | Physiological | FLo Err (FLOW ERROR) | FLo Err (FLOW ERROR) | 0
LVP 01 | MediumPriorityAlarm | Technical | AC POWER FAIL | AC POWER FAIL | 0
LVP 02 | HighPriorityAlarm | Technical | BATTERY EMPTY | BATTERY EMPTY | 1
LVP 03 | MediumPriorityAlarm | Technical | BATTERY LOW | BATTERY LOW | 1
LVP 04 | MediumPriorityAlarm | Physiological | TITRATION NOT CONFIRMED | TITRATION NOT CONFIRMED | 0
LVP 05 | MediumPriorityAlarm | Physiological | VTBI DONE | VTBI DONE | 0
LVP 06 | MediumPriorityAlarm | Technical | Attention (with " 3 beeps") | Attention (with " 3 beeps") | 0 | *
LVP 07 | HighPriorityAlarm | Physiological | DOWNSTREAM OCCLUSION | DOWNSTREAM OCCLUSION | 0
LVP 08 | HighPriorityAlarm | Physiological | AIR IN LINE | AIR IN LINE | 1
LVP 09 | HighPriorityAlarm | Physiological | AIR-IN-LINE | AIR-IN-LINE | 1
LVP 10 | HighPriorityAlarm | Technical | DOOR OPEN | DOOR OPEN | 0
LVP 11 | HighPriorityAlarm | Physiological | FLOW ERROR | FLOW ERROR | 0
LVP 12 | HighPriorityAlarm | Physiological | FREE FLOW | FREE FLOW | 0
LVP 13 | HighPriorityAlarm | Technical | FLOW SENSOR DISCONNECT | FLOW SENSOR DISCONNECT | 0
LVP 14 | MediumPriorityAlarm | Physiological | SET VTBI | SET VTBI | 0
LVP 15 | MediumPriorityAlarm | Physiological | SET MISLOAD | SET MISLOAD | 0
LVP 16 | HighPriorityAlarm | Technical | SAFETY CLAMP | SAFETY CLAMP | 0
LVP 17 | HighPriorityAlarm | Technical | DOOR CLOSE INCOMPLETE | DOOR CLOSE INCOMPLETE | 0
LVP 18 | MediumPriorityAlarm | Technical | SET NOT FITTED | SET NOT FITTED | 0
LVP 19 | HighPriorityAlarm | Physiological | UPSTREAM OCCLUSION | UPSTREAM OCCLUSION | 1
LVP 20 | HighPriorityAlarm | Technical | WRONG SET | WRONG SET | 0
LVP 21 | HighPriorityAlarm | Technical | AIR IN LINE | AIR IN LINE | 1
LVP 22 | MediumPriorityAlarm | Physiological | KVO Attention | KVO Attention | 0
LVP 23 | HighPriorityAlarm | Technical | LEVER OPEN | LEVER OPEN | 0
LVP 24 | HighPriorityAlarm | Technical | NO FLOW | NO FLOW | 0
LVP 25 | HighPriorityAlarm | Physiological | NO PRIMARY VTBI | NO PRIMARY VTBI | 0
LVP 26 | HighPriorityAlarm | Technical | PRIMARY FLOW DURING SECONDARY | PRIMARY FLOW DURING SECONDARY | 0
LVP 27 | MediumPriorityAlarm | Technical | SET CLOCK | SET CLOCK | 0
LVP 28 | MediumPriorityAlarm | Physiological | NEAR END OF INFUSION | NEAR END OF INFUSION | 0 | *
SE 01 | HighPriorityAlarm | Physiological | ACCUMULATED AIR IN LINE | ACCUMULATED AIR IN LINE | 0
SE 02 | HighPriorityAlarm | Physiological | AIR IN LINE | AIR IN LINE | 1
SE 03 | MediumPriorityAlarm | Technical | HOLD TIME EXCEEDED | HOLD TIME EXCEEDED | 0
SE 04 | MediumPriorityAlarm | Technical | Battery Low | Battery Low | 1
SE 05 | MediumPriorityAlarm | Technical | COMPUTER LINK FAILURE | COMPUTER LINK FAILURE | 0
SE 06 | HighPriorityAlarm | Physiological | Resistance Alert | Resistance Alert | 0
SE 07 | HighPriorityAlarm | Technical | FLOW SENSOR UNPLUGGED | FLOW SENSOR UNPLUGGED | 0
SE 08 | HighPriorityAlarm | Physiological | LATCH OPEN | LATCH OPEN | 0
SE 09 | HighPriorityAlarm | Physiological | OCCLUSION DOWNSTREAM | OCCLUSION DOWNSTREAM | 1
SE 10 | HighPriorityAlarm | Physiological | PRIMARY FLOW DETECTED DURING SECONDARY | PRIMARY FLOW DETECTED DURING SECONDARY | 0
SE 11 | HighPriorityAlarm | Physiological | NO UPSTREAM FLOW DETECTED | NO UPSTREAM FLOW DETECTED | 0
SE 12 | MediumPriorityAlarm | Physiological | SET OUT | SET OUT | 0
SE 13 | MediumPriorityAlarm | Technical | SETUP TIME EXCEEDED | SETUP TIME EXCEEDED | 1
SE 14 | HighPriorityAlarm | Technical | KEY STUCK | KEY STUCK | 0
SE 15 | HighPriorityAlarm | Physiological | OCCLUSION UPSTREAM | OCCLUSION UPSTREAM | 1
SE 16 | MediumPriorityAlarm | Physiological | VTBI = 0 | VTBI = 0 | 0
SE 17 | MediumPriorityAlarm | Technical | COMPUTER CONTROL RELEASED | COMPUTER CONTROL RELEASED | 0
SYR 01 | MediumPriorityAlarm | Technical | AC POWER FAIL | AC POWER FAIL | 0
SYR 02 | HighPriorityAlarm | Technical | DRIVE DISENGAGED | DRIVE DISENGAGED | 0
SYR 03 | HighPriorityAlarm | Technical | BATTERY EMPTY | BATTERY EMPTY | 1
SYR 04 | MediumPriorityAlarm | Physiological | END OF INFUSION | END OF INFUSION | 1
SYR 05 | MediumPriorityAlarm | Physiological | END OF INFUSION | END OF INFUSION | 1
SYR 06 | MediumPriorityAlarm | Physiological | END OF INFUSION | END OF INFUSION | 1
SYR 07 | MediumPriorityAlarm | Technical | BATTERY LOW | BATTERY LOW | 1
SYR 08 | MediumPriorityAlarm | Physiological | NEAR END OF INFUSION | NEAR END OF INFUSION | 0 | *
SYR 09 | HighPriorityAlarm | Physiological | OCCLUSION | OCCLUSION | 1
SYR 10 | HighPriorityAlarm | Technical | CHECK SYRINGE | CHECK SYRINGE | 1
SYR 11 | MediumPriorityAlarm | Technical | TITRATION NOT CONFIRMED | TITRATION NOT CONFIRMED | 0
SYR 12 | HighPriorityAlarm | Physiological | VTBI DONE | VTBI DONE | 0
SYR 13 | HighPriorityAlarm | Physiological | VTBI DONE | VTBI DONE | 0
SYR 14 | HighPriorityAlarm | Physiological | VTBI DONE - STOP | VTBI DONE - STOP | 0
SYR 15 | MediumPriorityAlarm | Technical | Attention (with " 3 beeps") | Attention (with " 3 beeps") | 0 | *
SYR 16 | HighPriorityAlarm | Technical | PRESSURE DISC OUT | PRESSURE DISC OUT | 0
SYR 17 | MediumPriorityAlarm | Technical | COMMUNICATIONS TIMEOUT | COMMUNICATIONS TIMEOUT | 0
SYR 18 | HighPriorityAlarm | Technical | INFUSION LINE DISCONNECTION | INFUSION LINE DISCONNECTION | 1
SYR 19 | HighPriorityAlarm | Physiological | EXCESSIVE RATE | EXCESSIVE RATE | 0
SYR 20 | HighPriorityAlarm | Technical | TCI TAG NOT RECOGNISED | TCI TAG NOT RECOGNISED | 0